
小鼠代谢笼检测系统操作流程
1.预热:接通代谢笼电源，打开配套电脑系统，进入软件Oxymax Windows V4.21界面，

预热时间为2小时 ，见下图1。

[image: image1.png]Oxymax Windows V4.21,COM1,C:\Program Files\OxymaxWinV421\Configl0B0003. ini % EH





2.设置:点击Experiment → Experiment File Open，选择080003文件，见下图2。
[image: image2.png]Experiment_Configuration Tools_Window About_Ext

Open Experiment File

Lookin: [ £ Config x| « @ ek

My Recent
Documents

@

Deskiop

=]

My Docurents

o Corpute
€ e [wom = Open
MoNawok  Flesolpe  [BPRes =] T
s

I~ Open as eadorly

4





3.气体校准:打开混合气钢瓶和氮气钢瓶的气体阀，调整压值使其都稳定在0.04-0.05之间，见下图3。

[image: image3.jpg]



点击Experiment → Calibrate 进行气体校准，见下图4。

[image: image4.png]@& Oxymax Windows V4.21,COM1

Propertiss

Event

Configuration

Feeding
Bouts

Activity

Temperature|

\Program Files\0xymaxWinV421\Config\0B000:

Event Log

Entor Log

| gow/cage

CHANNEL

DATE

Vo2

02007

D02

ACCOZ

VEoZ

Coan

CO20UT

DC0Z

[ACCCD?]

RER

HEAT

FLOW

mifkg/he

%

mifkg/he

B

%

L

keal/hr

PM





首先进行O2校准，点击O2  CALIBRATE → NEXT，见下图5。

[image: image5.png]@ Oxymax Windows V4.21,COM1,C:\Program Files\0xymaxWinV421\Config\0B000:

Experiment _ Configuration Tooks Window AboutExt

(® Oxymax Sensor Calibration

- Pump Velve =S ample Pump—— Enor Log
CoTestN || @ On
© Rel/caL D02 VEoZ CO20UT [Accoo?]

N2 ® i mivkg/hr % L

-Piessure Reaulator
& 0n

c of

~Sauce Pump:
&N

© o

rEsparsion teiface Valye Select—

345678

12
AR FLEY ccecccccc
El

LPM

&

e (@ 0xymax 02 Sensor Calibration,

1 Defaut_|

PRESSURE — | ———— | | |[0xymax 02 Sensor Offset Calibration
mmHg 02 [F02 i
o | T Ensure that the 02 offset gas is connected to the configured port. Ensure that the tank is

tumed on and the regulator is set to 5 psi. A small flow of gas should be apparent
overflowing from the open tee.

Group 1 02 offset channel = System Sample Pump N2 port
Group 1 02 span channel = System Sample Pus
(Click NEXT when ready.





轻微旋转 Fine Offset（ Low Gas Point）旋钮，使O2值稳定在0.00±0.01，数值变绿 → NEXT，见下图6。

[image: image6.png]- Pump Velve =S ample Pump—— Enor Log
CoTestN || @ On
© Rel/caL D02 VEoZ CO20UT [Accoo?]

N2 ® i mivkg/hr % L

-Piessure Reaulator
& 0n

c of

~Sauce Pump:
&N

© o

rEsparsion teiface Valye Select—

12345678
ccecccccc
El

AIR FLOW

LPM o

(& Oxymax 02 Sensor Calibration

PiEsse

+ Default_|
PRESSURE — | ———— | | [oxymax PARAMAGNETIC 02 Sensor Offset Calibration
mmHg 2 ooz

Using the COARSE and FINE offset controls on the front panel of the
CALGRATE | _CALIBRATE 9 ° ot b
PARAMAGNETIC 02 sensor(s), adjust for an 02 reading of :

Group 1 Offset=0.00

Hit NEXT to con

CANCEL | NEXT





轻微旋转 Span（ Increment From Low Gas Point ）旋钮，使O2值数值变绿 → NEXT → FINISH，见下图7，8。

[image: image7.png]- Pump Velve =S ample Pump—— Enor Log
CoTestN || @ On
& Rel/CAL D02 VEoZ CO20UT [Accoo?]

N2 ® i mivkg/hr % L

-Piessure Reaulator
& 0n

c of

~Sauce Pump:
&N

© o

rEsparsion teiface Valye Select—

12345678
AR FLEY ccecccccc
El

PiEsse

tDefault_|
PRESSURE =

mmmHg o2 oz Using the GAIN control(s) on the front panel of the
cauerate | | |02 sensor(s), adjust for a O2 reading of

Close Group 1 Span=20.90

Hit NEXT to cortinue

(& Oxymax 02 Sensor Calibration

CANCEL | NEXT





[image: image8.png]- Pump Velve =S ample Pump—— Enor Log
CoTestN || @ On
& Rel/CAL D02 VEoZ CO20UT [Accoo?]

N2 ® i mivkg/hr % L

-Piessure Reaulator
& 0n

c of

~Sauce Pump:
&N

© o

rEsparsion teiface Valye Select—

12345678
AR FLEY ccecccccc
El

(& Oxymax 02 Sensor Calibration

PiEsse

PRESSURE Dol |

= ——

mmmHg oz oz Calibration Complete:
_CaLBRATE |

Close Turn off calibration gas if finished with all sensor calibrations.

Hit FINISH to continue.





完成O2校准后进行CO2校准，点击CO2  CALIBRATE → NEXT →数值变绿 → FINISH，校正完成后关闭气阀，见下图9，10。

[image: image9.png]Entor Log
€ TestN || @ On
 Ref/CAL D02 VEoZ CO20UT [Accoo?]

p
N2 Gl mivkg/hr % L

@ on

0.317E8 oo

& on
© OFF

Oxymax CO2 Sensor, Calibration
AIR FLOW

LPM

Ensure that the CO2 zero gas is connected to the configured port. The zero gas can
bbe pure nitrogen or room air that has been scrubbed through a soda lime canister. I
nitrogen is being used, ensure that the tank is tuned on and the regulator is setto ~
PRESSURE __setDefait_| 5 psi. A smallflow of gas should be apparent overflowing from the open tee.

| B
mmmHg 02 co2 Ensure that the CO2 span gas is connected to the configured port. Ensure that the
_CALIBRATE | tank is turned on and the regulator is setto ~5 psi. A small flow of gas should be
apparent overflowing from the open tee.

(Group 1 CO2 offset channel = Systern Sample Pump N2 port v

oL | et





[image: image10.png]& Oxymax Windows V4.21,COM1

(® Oxymax Sensor Calibration

AIR FLOW
LPM

PRESSURE
mmHg

\Program Files\0xymaxWinV421\Config\0B000:

Experiment _ Configuration Tooks Window AboutExt

- Pump Valve —Sample Pump——
CoTestN || @ On

& Rel/CAL

ot c of

Entor Log

D02 VEoZ

CO20UT

[ACCCD?]

mifkg/he

%

L

-Piessure Reaulator
& 0n

c of

~Sauce Pump:
&N

© o

rEsparsion teiface Valye Select—

12345678
ccecccccc
El

&

PiEsse

__SetDefaut_|
= ——

oz C02

CALBRATE | [_CALIERATE

Oxymax CO2 Sensor Calibration

Oxymax CO2 Sensor Calibration
€02 Sensor calibration complete.

Click FINISH to save calibra
without saving constants.

Group 1 Offset=0.0000 volts  Gail

FINISH





4.实验设置:将已称重或做过核磁检测的小鼠放入代谢笼中，轻扣好盖子，勿蛮力 。点击Experiment → Setup，在Data File Name栏中输入数据保存路径及文件名，（保存路径:请以本软件根目录+20XX-XX文件夹+使用组别文件夹+使用者文件夹建立，文件名不能是中文） ，依次输入笼中小鼠的ID值，Mass值（可根据实验需求选用总体重或肌肉值），ENTER确认，（Flow值不要修改），如笼位未放满，可右键点击Grp/Cage栏去除。点击Apply → Close 完成设置，见下图11。

[image: image11.png]@ Oxymax Windows V4.21,COM1,C:\Program Files\OxymaxWinV421)\Configl0B0003. ini

Experiment _ Configuration Tooks Window AboutExt

@ Experiment

Feeding

T Activity | Temperature| Enor Log

Metabolic |  Feeding

INTERVAL CHANNEL DATE Vo2 D02 VEoZ CO20UT [ACCCD?]
] goup/cage mifkg/he mifkg/he % L

(& Experiment Setup [
Data i Name.
Browse.. | SetDefaut Path
Copy File Path
] Biowse.. | SetDefaul Pah | Clear

Use ENTER key to accept data enty. Cocle Timefmss|

Use ESC ke to exit edit mode. 2000 (2230)

Gip/Cage] _ SubedllD | Mass(gl [Flon((PM] | Gre/Cage 0101
o101 1 0 05 LABEL : INFORMATION
0102 05 [ I
0103 05 I
0104 05
0105] 05 [
0106 05 I

[
[
[
[

007| 05
0108] 05

e | oo | i it





5.实验开始:完成实验设置后点击Experiment → Run 开始实验，见下图12。

[image: image12.png]@& Oxymax Windows V4.21,COM1
3 Event

Setup
Calbrate

Propertiss

Configuration

ing

Feeding
Bouts

Activity

Temperature|

\Program Files\0xymaxWinV421\Config\0B000:

Event Log

Entor Log

CHANNEL

DATE

Vo2

02007

D02

ACCOZ

VEoZ

Coan

CO20UT

DC0Z

[ACCCD?]

RER

HEAT

FLOW

| gow/cage

mifkg/he

%

mifkg/he

B

%

L

keal/hr

PM





实验结束及数据导出:实验结束后点击Experiment → Stop，点击File → Create Files → CSV，导出数据，见下图13，14。

[image: image13.png]@ Oxymax Windows V4.21,COM1,C:\Program Files\0xymaxWinV421\Config\0B000:

4] Event Confiuration

Feeding
Bouts

Activity | Temperature| EventlLog | EmorLog

Refil Feeders CHANNEL DATE V07 | 02N | 02007 | D02 | ACCO2 | VEDZ | COPN [ COZ0UT | DCOZ [ACCCOZ] AER | HEAT | FLOW
Propertes oup/cage kg | % % [ T % L calth | TP





[image: image14.png]@& Oxymax Windows V4.21,COM1

Experimert_Event Configuration

oeding
Bouts

Activity

Temperature|

\Program Files\0xymaxWinV421\Config\0B000:

Event Log

Entor Log

INTERVAL _Steep.

INEL

DATE

Vo2

02007

D02

ACCOZ

VEoZ

Coan

CO20UT

DC0Z

[ACCCD?]

RER

HEAT

FLOW

]

Goup/cage

mifkg/he

%

mifkg/he

B

%

L

keal/hr

PM





7.系统清洁:将小鼠取出后，将使用后的笼子、水瓶、配件用清水洗净，晾干待用。
注意事项: 1. 仪器长期运行，2周校准气体1次即可。

          2. 仪器有工作异常、损坏，配件丢失请及时通知管理员。

          3. 除氨剂变色后，请及时通知管理员更换。

          4. 干燥剂由蓝色变为紫红色，请及时通知管理员更换，频率随空气湿度而定。



